| Inspected by | TJF | Date | 4/3/2026 |
| --- | --- | --- | --- |

Sensor Information

| ECO S/N | BBFL2WB-1755 |  |
| --- | --- | --- |
| RMA | 00001683 |  |
| Types | ☐ Real time (-RT) or base ☐ Battery (-B) ☐ Wiper/ Shutter (-W or -S) ☒ Battery AND Wiper/ Shutter (-SB or -WB) | ☐ -AP2/ -RT2K ☐ Deep (-D/-RT6K/ -6K) ☐ -SLC ☐ Puck |
| Number of Channel | 3 |  |
| Channel 1 | NTU |  |
| Channel 2 (if applicable) | CHL |  |
| Channel 3 (if applicable) | CDOM |  |

Hardware/ Parts

| RT, Base, Shutter unit | Shutter/ Wiper unit | Battery unit | -SLC unit |
| --- | --- | --- | --- |
| ☐ Mounting bracket | ☒ Copper face plate ☒ Copper Shutter/ Wiper | ☒Dummy battery plug/ collar | ☐ Copper face plate ☐ Molex cover |
| -AP2 unit | -RT2K unit | Deep unit | CDOM unit only |
| ☐ Anode ☐ Saddle | ☐ Anode | ☐Max 6000m label | ☒ UV warning label |

| ☒ | Lock collar, dummy plug/ collar |
| --- | --- |
| ☒ | Protective end cap has same SN as the unit |
| ☒ | Serial Number Label |

Physical Inspection

| ☒ | Sensor SN listed above matches with the label on the unit |
| --- | --- |
| ☒ | Fishing line is present in slot |
| ☒ | Connector pins on bulkhead are not corroded |
| ☒ | Optics window has no crack, obvious debris, |
| ☒ | Wiper/ Shutter unit: Rubber blade(s) from wiper/ shutter are contacting the optics surface |
| ☒ | Bulkhead tightened |

Documentation

| USB | ☒ Check that USB folder’s name includes unit SN |  |  |  |
| --- | --- | --- | --- | --- |
| Manuals | ☒ Sensor user manual ☒ Triplet-W Manual ☒ ECOView software ☒ ECOView user manual |  |  |  |
| Char Sheets |  | Channel 1 | Channel 2 | Channel 3 |
|  | Date | ☒ | ☒ | ☒ |
|  | S/N | ☒ | ☒ | ☒ |
|  | Temp | ☒ | ☒ | ☒ |
|  | Dark Counts | 50 | 49 | 50 |
|  | Dark Count (40-70) (CDOM: 40-80) | ☒ | ☒ | ☒ |
|  | Analog Darks (V) | N/A | N/A | N/A |
|  | Scale Factor != 1 | ☒ | ☒ | ☒ |
|  | Saturation | 4130 | 4130 | 4130 |
|  | ☐ Char Sheet for Post-cruise and/or Secondary calibration (if applicable) |  |  |  |

Documentation (Continued)

| Device File | ☒ Device file present and labeled with sensor SN ☐ Single Channel FL contain both: ☐ Dev File ☐ ENGR Dev File ☐ AP2 only: use OEM dev file ☐ 2 Channel = 5 blank columns in dev file ☐ 3 Channel = 7 blank columns in dev file ☐ SLC & Puck & RT2K only: ☐ 2 Channel = 7 columns in dev file ☒ 3 Channel = 9 columns in dev file |
| --- | --- |
| Triplet-W only | ☒ “TripletW.zip” file ☒ “Triplet-W setup commands” |
| Repair Sheet | ☒ Repair sheet ☒ Correct S/N in Repair Sheet |
| Final USB Check | ☒ Verify no other files are in USB Folder ☒ QA Checklist IS NOT in USB Folder |

Functionality (Open EcoView)

| Saturation in EcoView | 4130 | 4130 | 4130 |
| --- | --- | --- | --- |
| Saturation equal to Char Sheet | ☒ | ☒ | ☒ |
| Dark AVG from “D” file | 50 | 48.98 | 50 |
| Dark Value = Char Sheet | ☒ | ☒ | ☒ |

| ☒ | Verify Saturation: - 16380 for 1ch FL / 1 ch NTU units - 4130 for 2ch, 3ch, and 1ch BB |
| --- | --- |
| ☒ | Verify that serial number shown in EcoView output is correct |
| ☒ | Return Settings to Incoming Settings |
| ☐ | Clear memory (if applicable) $emc |
| ☒ | Internal battery voltage (8.5V min) |

Functionality (Open EcoView) (Continued)

| 1Ch Chlorophyll unit only: | ☐ |  | Use $par and verify that ugl matches scale factor and off matches dark counts |
| --- | --- | --- | --- |
| FLBBFL-/ TripletW units only | ☒ |  | Use $par and verify that offset M1d, M2d, M3d in unit menu is as written in dev file |
|  | ☒ |  | Use $par and verify that scale factor M1s, M2s, M3s in unit menu is as written in dev file |
| CDOM unit only |  | ☒ | Passing result for CDOM test (in main unit folder) |

Oracle
SBE Service Depot Repair > Repair Orders > Enter RMA > Select Instrument

| ☒ | Verify Repair Sheet content matches Debrief Notes |
| --- | --- |
| ☒ | Verify Repair Sheet content matches customer request in the “Preliminary Service Quote Request” email in the Job Folder |
| ☒ | Verify correct S/N on Repair Sheet |
| ☒ | Verify “Material Consumed” matches Repair Sheet |
| ☒ | Place USB folder into Job RMA Folder (linked in Oracle) |
| ☒ | Change status to Done in Oracle |
| ☒ | Enter in Description line: “PASSED QA” – INITIALS- DATE |

Administrative

| ☒ | Update info into QA Log Book |
| --- | --- |

Screen shot of the EcoView Raw Data

Screen shot of the EcoView Setup Tabs