ECO SERVICE TRIAGE SHEET

| Name: LP | Triage Date: 4/1/2026 |
| --- | --- |

| Customer Name: NOAA PMEL\Pacific Marine Environmental Lab\UW\JISAO | RMA: 00001683 |  |
| --- | --- | --- |
| SN: BBFL2WB-1755 | Firmware Version: Ver Triplet5.33 |  |
| Build date: 10/31/2018 | Last-serviced date: 1/28/2021 | ☐ Warranty? |

(Warranty policy: 2 years for ECO production, 1 years for ECO service)

| Summary |  | ☐ Post-deployment Cal. | ☐ Secondary Cal. | 1 Star Priority |
| --- | --- | --- | --- | --- |
| Customer Notes/ Evaluation: Triage: No problems were found at this time. Std srv. [4-1 LP] |  |  |  |  |
| Reminder | - Record pkt 40 log for all channels (except backscatter) and save to Incoming folder. - Transfer part from B1 to BSV when removing part from physical location. - Complete Service Completion Checklist before delivering the unit to QA. |  |  |  |

Admin

| ☒ | Verify that the SN written on cap matches the SN of the unit label and paperwork |
| --- | --- |
| ☒ | Create a new calibration folder “CAL#” |
| ☒ | Create Incoming Folder inside calibration folder |

Functionality

| ☐ | Inspect optical head for cracks and delamination |
| --- | --- |
| ☒ | Supply power to sensor; Record current drawn: 32.16 mA |
| ☒ | *Non-CDOM sensors only: make sure the LEDs light. (UV LEDs may cause retina damage) |
| ☒ | Save $mnu and $par settings and 10 samples in Incoming folder |
| ☐ | ASV Setting (Single channel Instruments): Enter value |
| ☒ | Capture memory from instrument if applicable. (N/A for RT and puck sensors) |
| ☒ | Record Saturation for each channel: So1 + sat = 4130 So2 + sat = 4130 So3 + sat = 4130 |
| ☒ | Verify unit is saturating at above recorded saturation values |
| ☐ | Verify analog saturation is between 4.90-5.01 V (if applicable) |
| ☒ | Verify shutter operation. $mvs1/mvs 0 commands. Issue found: NONE |
| ☒ | Triplet-W Only: Verify new wiper motor design. If old motor exists, it must be replaced. |

Functionality Contd.

| ☒ | Record sleep current (-W, -S, & Memory units): 119.2 µA |
| --- | --- |
| ☒ | Perform CDOM test and save result to main unit folder. **CDOM test only needs to be performed ONCE in an optics head lifetime.** |
| ☐ | If unit is unserviceable or failed CDOM test, place on HOLD shelf, update notes in Oracle, and email Lead and Admins with details. (see “Replacement – Warranty .docx” for template) |

Calibration

|  | Solution/ Wavelength | Range |
| --- | --- | --- |
| Channel 1 | NTU (If backscatter, **** nm) | 200 |
| Channel 2 | CHL (If backscatter, **** nm) | 75 |
| Channel 3 | CDOM/ PTSA (If backscatter, **** nm) | 375 |

Oracle

| ☒ | Check Customer Notes (post-deployment calibration request, secondary calibration, functionality issues, special request, etc.) and enter details into Triage sheet. |
| --- | --- |
| ☒ | Check Expense tab for Secondary Calibration (SRVSECCAL) or other specified service(s). |
| ☒ | Enter spare parts that will be replaced (faceplate, wiper, etc.) |
| ☒ | Update Status Field |
| ☒ | Update Description Field: “Check in completed [Date / Initials]”. |
| ☒ | Save. |

| COMMONLY REPLACED PARTS (see “ECO Service Part Numbers.xls” for an extended list) |  |  |  |
| --- | --- | --- | --- |
| Shutter face plate | MAA-220524 | 2CHN SLC face insulator | MAA-230623 |
| Shutter | BAA-000017 | 3CHN SLC face insulator | MAA-230681 |
| Shaft seal | MAA-220527 | FLBBCD-SLC face plate | MAA-230613 |
| 9V Battery | EXA-ZX0064 | BBFL2-SLC face plate | MAA-230671 |
| TripW face plate | MAA-230330 | FLBB-SLC face plate | MAA-230615 |
| TripW wiper | BAA-251621 | PAR Bio-Wiper | BAA-220484 |
| TripW motor | BAA-537892 |  |  |